Manual Blood Pressure Trainer

Product and Programme Plan, Version 4

Learning architecture, curriculum, media asset production, and the way forward.

| Field | Value |
| --- | --- |
| Document identifier | PPP-MBPT-004 |
| Version | 4.0 draft, for internal review |
| Date | August 9, 2026 |
| Author | Melissa Tully, Founder, Chief Executive Officer and Learning Engineer |
| Status | Draft. Standalone. Reads without reference to any prior revision. |
| Companion workbook | Encountive MBPT UX and Curriculum Crosswalk, which holds the activity level detail |
| Correction carried | The objective set contains 39 enabling objectives, not 32. Earlier documents undercounted. See section 2.2. |

# Contents

# 1. The product

## 1.1 What it is

A training system for the auscultatory method of manual blood pressure measurement, for pre-licensure and entry level healthcare cohorts, initially patient care technician and practical nursing programmes.

The learner keeps real equipment: a real cuff, a real bulb and valve, a real aneroid gauge, a real acoustic stethoscope. The software supplies what the classroom cannot: a circulation with a chosen pathology, an instructor who is watching every attempt, and a record of what actually happened.

## 1.2 The three problems it solves

- Instructor time does not scale. One instructor supervising twenty learners cannot observe each learner's cuff placement, deflation rate and sound identification. Feedback arrives late, in aggregate, or not at all.

- Human variation is unavailable. A cohort practises on each other. Healthy young adults of similar size produce a narrow band of normal readings, so learners graduate having never heard an auscultatory gap, never measured an arm needing a large adult cuff, and never encountered an irregular rhythm.

- Remote and hybrid learners have no equivalent practice. Distance cohorts either travel to campus or practise unsupervised with no feedback loop.

## 1.3 How it works, in one paragraph

The learner performs the procedure on a partner with real equipment. A camera reads the needle on the real aneroid dial, so the system knows the cuff pressure the learner is actually producing. Recorded Korotkoff sounds from a real patient with the chosen condition are played into the learner's earphones, indexed to that measured pressure, so the sounds appear and disappear at the correct pressures no matter how fast or slowly the learner deflates. The learner marks systolic and diastolic, states an interpretation and an action, and is scored deterministically against expert established ground truth.

Scope boundary, stated plainly This is preparation and retention practice. It does not replace supervised measurement on live human beings, and no claim to the contrary should appear in any product or marketing material. Its defensible claim is that learners arrive at live practice having already achieved procedural fluency and having already heard variation they would otherwise never encounter.

## 1.4 The three things this document plans

- The learning architecture: what a competent learner must be able to do, and how each of those things is observed.

- The curriculum and its crosswalk: the activities that teach and assess each objective, with automatic checks that nothing is taught but unscored, or scored but never taught.

- The media asset production programme: how the clinical recordings that the whole product depends on actually get made. This is the longest lead time item in the programme and it has never had a plan of its own.

# 2. Learning architecture

## 2.1 Terminal objective

TO-1 Given a patient of any body habitus, a calibrated aneroid sphygmomanometer, and a stethoscope, the learner will obtain and interpret an auscultatory blood pressure measurement that falls within plus or minus 4 millimetres of mercury of the reference value on both systolic and diastolic pressure, using correct technique at every scored step, and will state the appropriate follow up action for the value obtained.

Mastery requires the terminal objective on three consecutive attempts, across at least two body habitus categories and at least two clinical cases, with no scaffolding present. A single successful attempt is a data point that could have been produced by rounding luck, which is why rounding bias is monitored as a trend rather than per attempt.

The four millimetre criterion is not arbitrary. It matches the scoring tolerance, and it matches the inter-rater agreement threshold used to accept a case into the library. A case whose three expert raters cannot agree within four is rejected rather than averaged, so scoring a learner more tightly than the experts achieved would be indefensible.

## 2.2 The objective set

Thirty nine enabling objectives across six domains. Earlier drafts of this programme stated thirty two, which was an arithmetic error carried forward without checking. The correct counts are seven, six, six, eight, eight and four.

| Domain | Title | Count | What it covers |
| --- | --- | --- | --- |
| A | Preparation and patient safety | 7 | Contraindication screening, positioning, rest, patient explanation, hygiene |
| B | Cuff selection across body habitus | 6 | Measuring, selecting, consequence of mis-sizing, application, boundary judgement |
| C | Inflation and deflation control | 6 | Palpated estimate, inflation adequacy, deflation rate, re-inflation discipline |
| D | Auscultation and identification | 8 | Stethoscope placement, phase identification, gap detection, precision reading |
| E | Interpretation and action | 8 | Classification, escalation, documentation, reliability and repeat |
| F | Self monitoring and practice regulation | 4 | Trend awareness, error attribution, practice planning |

### Domain A, preparation and patient safety

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| A1 | Screen the limb for contraindications and select an alternative site or escalate, identifying every contraindicated limb | Cognitive, Apply | Gate response |
| A2 | Position the patient with back supported, feet flat, legs uncrossed and the arm supported at heart level, all four before inflation | Psychomotor, Guided response | Geometric or vision check |
| A3 | State how each positioning fault changes the reading and in which direction | Cognitive, Understand | Spoken or typed response |
| A4 | Allow five minutes of quiet rest before inflation, with no talking | Psychomotor, Set | Timer, steps completed |
| A5 | Identify conditions that invalidate a reading and defer or document accordingly | Cognitive, Apply | Gate response |
| A6 | Explain the procedure to the patient and obtain cooperation before touching the limb | Affective, Responding | Speech event detection, partial |
| A7 | Perform hand hygiene before patient contact | Psychomotor, Set | Self report only, never a measured competency |

### Domain B, cuff selection across body habitus

This domain carries the human variation acceptance criterion. It is the reason the product exists.

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| B1 | Measure arm circumference at the midpoint between acromion and olecranon, to within one centimetre of reference | Psychomotor, Guided response | Entered value, independently checked |
| B2 | Select the correct cuff from a measured circumference, with no error across ten trials spanning four size categories | Cognitive, Apply | Cuff selected against form detected |
| B3 | Predict the direction and approximate magnitude of error from an undersized and an oversized cuff | Cognitive, Analyze | Response, plus consequence simulation |
| B4 | Apply the cuff on bare skin, bladder centred over the brachial artery, lower edge 2 to 3 centimetres above the fossa, index line verified | Psychomotor, Guided response | Vision or spatial check |
| B5 | Given a boundary circumference or a markedly conical arm, justify the cuff chosen and state the residual uncertainty | Cognitive, Evaluate | Spoken justification |
| B6 | Recognise when no available cuff fits and state the correct alternative rather than proceeding | Cognitive, Evaluate | Gate response |

### Domain C, inflation and deflation control

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| C1 | Palpate the radial pulse during inflation and record its disappearance within 10 mmHg of auscultated systolic | Psychomotor, Guided response | Pressure trace, timed hold |
| C2 | Inflate to at least 30 mmHg above the palpated estimate, on every attempt | Cognitive, Apply | Pressure trace peak |
| C3 | Control deflation at 2 to 3 mmHg per second for the full descent with no excursion outside the band | Psychomotor, Complex overt response | Pressure trace rate and excursions |
| C4 | Explain why underinflation risks reading systolic below a gap, and why over rapid deflation biases both values | Cognitive, Analyze | Response |
| C5 | Deflate fully and wait before any repeat inflation, never re-inflating on a partially deflated cuff | Psychomotor, Mechanism | Pressure trace re-inflation events |
| C6 | Detect an out of band deflation rate during the attempt and correct it without restarting | Cognitive, Evaluate | Pressure trace, trend |

### Domain D, auscultation and identification

The perceptual core, and the reason the clinical recordings in section 4 cannot be synthesised.

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| D1 | Place the stethoscope over the brachial artery, below the cuff edge, with full skin contact and no tubing contact | Psychomotor, Guided response | Vision or spatial track, partial |
| D2 | Identify phase one onset and mark systolic within 4 mmHg of reference | Psychomotor, Perception | Marks |
| D3 | Identify phase five disappearance and mark diastolic within 4 mmHg of reference | Psychomotor, Perception | Marks |
| D4 | Detect an auscultatory gap, state its pressure range, and explain how underinflation would falsify systolic | Cognitive, Analyze | Marks, response |
| D5 | Discriminate phase four muffling from phase five disappearance and state the populations in which phase four is recorded | Psychomotor and Cognitive | Marks, response |
| D6 | Read the needle to the nearest 2 mmHg without rounding to 5 or 10, across at least ten attempts | Psychomotor, Perception | Terminal digit trend |
| D7 | Recognise an irregular rhythm during auscultation and state that a single reading is unreliable | Cognitive, Analyze | Response |
| D8 | Identify inaudible or ambiguous sounds and state the corrective actions available | Cognitive, Evaluate | Response |

### Domain E, interpretation and action

The domain most training software omits. The clinically consequential failures happen after the number, not during it.

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| E1 | Classify a measured value against the current categories and state the classification | Cognitive, Apply | Response |
| E2 | State that a single elevated reading does not establish a diagnosis, and state the repeat requirement | Cognitive, Understand | Response |
| E3 | Given an irregular rhythm, take multiple readings, average them, and justify why | Cognitive, Apply | Marks across attempts, response |
| E4 | Identify a hypertensive crisis value and state the escalation action within scope, unprompted | Cognitive, Evaluate | Response with prompted flag |
| E5 | Identify a hypotensive value, relate it to the presentation, and state the action | Cognitive, Evaluate | Response |
| E6 | Document arm used, patient position, cuff size and time, omitting nothing | Cognitive, Apply | Response |
| E7 | Recognise a reading implausible against the palpated estimate or a prior value and repeat rather than record it | Cognitive, Evaluate | Marks, trace, response |
| E8 | State the correct action when the value differs materially from a documented prior value | Cognitive, Evaluate | Response |

### Domain F, self monitoring and practice regulation

Only assessable across attempts, which is something an instructor watching a single performance cannot do. This is the domain that justifies building a tutoring system rather than a checklist.

| ID | Objective | Level | Evidence |
| --- | --- | --- | --- |
| F1 | Identify a rounding bias in the learner's own values across at least ten attempts and state a corrective strategy | Cognitive, Evaluate | Terminal digit trend |
| F2 | Given an out of tolerance result, identify which step of the learner's own procedure most likely produced it | Cognitive, Analyze | Attempt record plus response |
| F3 | Select the next practice condition from the learner's own performance pattern and justify the choice | Cognitive, Evaluate | Selection plus mastery flags |
| F4 | Distinguish a technique error from a perceptual error in the learner's own performance | Cognitive, Analyze | Scorer output plus response |

## 2.3 Safety gates, three kinds

Prior planning treated safety as one concept. It is three, and they behave differently in code.

| Gate | Fires | Behaviour | Example |
| --- | --- | --- | --- |
| Contraindication | Before the procedure | Refuses the limb. The correct action is to select the other arm or escalate, not to proceed carefully. | Arteriovenous fistula, lymphoedema or prior axillary node dissection, intravenous line, injury |
| Technique | During the procedure | Blocks progression until satisfied. Repeatable within one attempt. | Cuff size, arm at heart level, five minute rest, inflation 30 above estimate, full deflation before re-inflation |
| Interpretation | After the measurement | Requires a stated action, not just a number. | Crisis value recognition, single reading does not diagnose, irregular rhythm requires averaging |

# 3. Curriculum and crosswalk

## 3.1 Structure

Six modules, twenty two lessons, sixty one activities, two hundred and ninety five minutes of designed instruction. Fifty two activities and two hundred and thirty nine minutes are available in the first release; nine activities and fifty six minutes depend on clinical recordings that do not yet exist.

| Module | Title | Lessons | Activities | Minutes | Domains |
| --- | --- | --- | --- | --- | --- |
| M0 | Orientation, equipment and setup | 1 | 3 | 8 | Setup, no objectives |
| M1 | Preparation and the procedure | 4 | 10 | 54 | A |
| M2 | Cuff selection across body habitus | 5 | 15 | 69 | B |
| M3 | Auscultation | 6 | 21 | 110 | C, D, F |
| M4 | Interpretation and action | 3 | 8 | 35 | E |
| M5 | Debrief and self regulation | 3 | 4 | 19 | F |

## 3.2 Four activity types

| Type | Count | Function | Feedback timing |
| --- | --- | --- | --- |
| Present | 7 | Teaches. No performance demanded, no scoring. Used sparingly. | Not applicable |
| Practice | 27 | Rehearses with feedback. Scaffolds visible. Does not count toward mastery. | Immediate |
| Assess | 24 | Scores. Only unaided scaffold states count toward mastery. | Held until the attempt ends |
| Reflect | 3 | Reviews across attempts. Never pass or fail. Owns the trend measures. | Presentational only |

The feedback timing rule matters more than it looks. Mixing immediate feedback into assessment teaches learners to fish for hints rather than to commit to a judgement, which is precisely the habit that fails at a bedside.

## 3.3 The scaffold ladder

The same objective can be met under four different conditions. Only one counts.

| State | What is present | Purpose |
| --- | --- | --- |
| Guided | The aid is visible during the performance. Overlay, numeric readout, rate pacer. | Builds the association between action and result. |
| Confirmatory | The aid appears only after the learner commits. | Teaches self checking rather than following. |
| Sound only | The gauge is occluded so the learner cannot substitute looking for listening. | Isolates the perceptual channel actually being trained. |
| Unaided | Nothing. Real equipment, real patient, no assistance of any kind. | The only state counting toward mastery. |

Because mastery depends on the scaffold state, every attempt record must carry which state it was performed under. Without that field the mastery rule cannot be evaluated and scaffolded success would be silently counted as competence.

## 3.4 What the crosswalk is for

The crosswalk maps every activity to the objectives it serves, and then checks the mapping in both directions. It is a live workbook rather than a table in this document, because it has to be recalculated every time an activity is added or a duration changes.

| Check | Question it answers | Current state |
| --- | --- | --- |
| Not covered | Is any objective served by no activity at all? | 0 |
| Taught but never scored | Is any objective practised but never assessed? | 0 |
| Scored without practice | Is any objective assessed with no teaching activity before it? | 0 |
| Release availability | Which objectives are unavailable until the clinical recordings exist? | 5 |
| Fully served in the first release | Which objectives are complete now? | 34 |

Why this check earns its place Run against the first draft of the curriculum, the crosswalk found twenty three faults: thirteen objectives were taught but never scored, and ten were assessed with no teaching activity before them. Those are exactly the failures that get built and only surface when a programme director asks how a competency is evidenced. Closing them produced two activities the draft was missing entirely, including a full end to end unaided attempt.

## 3.5 The five objectives held back

Five objectives cannot be delivered until the corresponding clinical recordings exist. They are named here rather than discovered later, and they are the entire justification for the production programme in section 4.

| Objective | Requires | Activities blocked |
| --- | --- | --- |
| D4, auscultatory gap detection | A recording containing a genuine auscultatory gap | A-3.6.1, A-3.6.5 |
| D5, phase four discrimination | A recording with a clear muffling point | A-3.6.2, A-3.6.5 |
| D7, irregular rhythm recognition | An atrial fibrillation recording | A-3.6.3, A-3.6.5 |
| E3, averaging in irregular rhythm | The same atrial fibrillation recording | A-4.3.2 |
| E5, hypotension recognition and action | A hypotensive recording with a patient presentation | A-4.2.2 |

The honest consequence: the first release covers preparation, cuff selection, inflation and deflation control, core auscultation, interpretation of ordinary values, and the whole self regulation domain. It does not cover pathology recognition. That is a clean, bounded claim and it is far better stated up front than discovered by a programme director in week three.

## 3.6 A session that fits a real timetable

A forty five minute rotation, four learners, one station. Designed so that nobody is idle and the equipment is never the bottleneck.

| Minutes | Learner at the station | The other three | Partner role |
| --- | --- | --- | --- |
| 0 to 5 | Orientation and equipment check, gauge zero | Same, on their own devices | Observes |
| 5 to 20 | Deflation rate control and marking practice | Rotate through the same | Holds the device, presents the arm |
| 20 to 33 | Scored checkpoint: measure, select, apply, place, measure | Continue practice | Presents the arm and a brief history |
| 33 to 40 | Interpretation and action, spoken | Continue practice | Observes |
| 40 to 45 | Debrief, error attribution, next practice chosen | Handover | Rotates |

Two design consequences fall out of this and neither is obvious from a feature list. The partner holding the device becomes an observer with a defined role rather than a passive arm, and observation is itself a documented learning mechanism. And handover has to complete in under sixty seconds or the rotation collapses, which makes handover a designed step rather than an interruption.

# 4. Media asset production

This is the longest lead time item in the programme, the one no amount of engineering shortens, and the one that has never had a plan of its own. It gets one here.

## 4.1 Four classes of asset

| Class | Asset | Quantity for release one | Why it cannot be synthesised |
| --- | --- | --- | --- |
| A | Clinical case recordings: aneroid gauge video, synchronised Korotkoff audio, and a reference pressure log | 14 accepted cases | Real Korotkoff sounds carry the timing irregularity, spectral character and subject to subject variation that define the perceptual objective. A tone generator produces a sound that trains a learner to recognise a tone generator. |
| B | Technique demonstration video: landmark location, cuff application, stethoscope placement, palpated estimate | 8 clips, 30 to 90 seconds each | Could be animated, but a real hand on a real arm shows the pressure, angle and hesitation that animation smooths away. |
| C | Habitus reference imagery: arms across the circumference range, with verified measurements | 12 arms across five categories | Needed to present size at scale for learners who cannot access the extremes in their cohort. |
| D | Instructor onboarding video: running a session, reading the reports, handling failure modes | 4 clips | Pilots fail on instructor confusion far more often than on software defects. |

Classes B, C and D are ordinary production work that can be done in a single day with a colleague and a tripod. Class A is the programme. Everything below concerns Class A unless stated.

## 4.2 The capture rig

One rig, built once, used for every session. Consistency across the library matters more than absolute quality, because a library captured under varying conditions is internally incomparable and no learner can be scored consistently across it.

| Element | Specification | Why this specifically |
| --- | --- | --- |
| Camera | Any phone or mirrorless capable of 1080p at 60 frames per second, on a fixed mount | Sixty frames gives finer needle sampling for the authoring pass. The camera never moves during a take. |
| Exposure, focus, white balance | LOCKED MANUALLY. No auto anything. | The single most important line in this table. Auto exposure hunting between frames changes needle contrast and ruins automated needle reading. |
| Framing | Dial face parallel to the sensor plane, filling roughly sixty percent of frame | Off axis framing introduces perspective distortion into the angle to pressure mapping. |
| Lighting | Two diffuse sources at roughly forty five degrees. Polarising filter if the dial glass shows a hotspot. | A specular reflection on the crystal can be mistaken for the needle or can hide it. |
| Audio | One digital stethoscope. One mode. One gain setting. Recorded to its own file at 16 kHz or above. | Different modes apply different filtering. Mixing them makes cases incomparable and is unrecoverable after the fact. |
| Reference pressure | An inline digital pressure sensor logging at 50 Hz or better to a laptop | Authoring only, one rig, never shipped to a customer. It provides ground truth pressure so the pipeline does not have to derive it from the needle, and it validates the needle reader at the same time. |
| Synchronisation | An electronic slate: a small LED in frame lit simultaneously with an audible tone into the stethoscope channel. A firm tap on the diaphragm is the backup. | Audio and video are separate files. Without a hard sync marker, every case carries an unknown offset and beat pressures are wrong by that offset. |

## 4.3 Take protocol

Identical for every take, on a printed card at the rig. Any deviation is noted on the take log rather than corrected silently.

- Slate: LED and tone, three seconds of stillness afterward.

- Subject seated five minutes, back supported, feet flat, arm supported at heart level, no talking.

- Correct cuff selected by measured circumference. Circumference recorded on the take log.

- Palpated estimate taken and recorded.

- Inflate to the palpated estimate plus 30 millimetres of mercury.

- Deflate at 2 to 3 millimetres of mercury per second through to below 40, monitored against the reference sensor.

- Full deflation and sixty seconds of rest.

- Repeat for three takes. Alternate arms if a fourth is needed.

A capture session per subject runs about forty five minutes including consent, setup and the rest intervals. Six subjects fit comfortably in a day. The capture itself is not the constraint.

## 4.4 Subjects and clinical variation

Priority order, with what each unlocks. The first four are what stands between the first release and pathology coverage.

| Priority | Subject condition | Unlocks | Recruitment difficulty |
| --- | --- | --- | --- |
| 1 | Atrial fibrillation, rate controlled | D7, E3, and two activities. The highest value single recording in the programme. | Moderate. Common in community volunteers over seventy. |
| 2 | Auscultatory gap, typically an older subject with stiff vessels | D4, and the trap that teaches why underinflation matters. | Hard. Cannot be recruited by diagnosis; must be found by screening candidates. |
| 3 | Hypotension, symptomatic or orthostatic | E5 | Moderate. |
| 4 | Clear phase four muffling, commonly pregnancy or a high output state | D5 | Moderate, but pregnancy adds consent complexity. |
| 5 | Stage two hypertension | Broader case variety at scale | Easy. |
| 6 | Large arm requiring a large adult or thigh cuff | Reinforces Domain B with real audio at the extreme | Easy. |
| 7 | Paediatric, school age | Paediatric variation | Hardest. Parental permission plus child assent. Longest lead time in the programme. |
| 8 | Wide pulse pressure, typically elderly | Case variety | Easy. |

The recruitment insight that shortens the schedule These do not have to be hospital patients. Community volunteers with a known diagnosis are far easier to recruit, far easier to consent, and produce identical recordings. Atrial fibrillation, hypertension and orthostatic hypotension are common in community settings. Approaching a senior centre, a cardiac rehabilitation programme or a church group is a completely different undertaking from approaching a hospital ward, and it removes an entire layer of institutional gatekeeping.

## 4.5 Consent and permissions

Flagged clearly: this is not legal advice and every point below should be confirmed with counsel and with the host institution before any recording occurs.

### The distinction that decides the timeline

Recording volunteers to produce an educational product is generally not human subjects research, because it is not a systematic investigation designed to develop generalisable knowledge. It is media production. Many institutions will still want a formal determination that it is not research, and that determination is typically a matter of weeks rather than the months a full review board protocol takes.

Evaluating whether the product improves learning is research, and that does require review board approval. These are two separate tracks with two separate timelines, and conflating them is what makes the media programme look like a six month problem when the production half of it is not.

| Track | What it is | Approval needed | Indicative timeline |
| --- | --- | --- | --- |
| Production | Recording volunteers to create teaching assets | Institutional permission, informed consent, media release. Likely a not human subjects research determination. | Weeks |
| Evaluation | Measuring whether learners trained on the product perform better | Full institutional review board protocol | Months. Start separately and later. |

### De-identification, which is easier here than it looks

The asset the product needs contains a dial, a cuff, a forearm, and heart sounds. It does not need a face, a voice, a name, or a date of service. Framed correctly, a Class A recording is inherently de-identified at the point of capture rather than in post production.

- Frame the dial and the cuff only. No face, no distinguishing marks, no tattoos or jewellery in shot.

- No speech on the stethoscope channel. The operator uses hand signals during a take.

- Subject identifiers live only on the take log, which is kept separately and is not shipped.

- Case identifiers are sequential and carry no clinical or personal reference.

### Consent must cover the actual use

A generic clinical photography consent is not sufficient. The form needs to name commercial educational distribution, indefinite retention, and the fact that the recording will be used to train learners at institutions other than the one where it was captured. A volunteer who consented to teaching use at their own clinic has not consented to a licensed product.

## 4.6 Production pipeline

| Stage | Work | Owner | Duration |
| --- | --- | --- | --- |
| P1 | Build the rig, freeze the capture protocol, write the take log template and the consent form | Melissa Tully with counsel | 2 weeks |
| P2 | Pilot capture: two willing colleagues, full protocol, end to end through to an accepted case | Melissa and Doug Tully | 1 week |
| P3 | Site identification and permission, consent form approval, not human subjects research determination | Melissa Tully with Nisha Patel | 3 to 6 weeks, runs in parallel from day one |
| P4 | Capture sessions. Six subjects per day. Two days covers priorities one to six. | Melissa Tully | 2 days plus scheduling |
| P5 | Post production: sync against the slate, trim to the take, verify no exposure drift, extract the pressure log | Contract or in house | 1 hour per accepted take |
| P6 | Annotation: needle read across the recording, beat onsets detected and tagged with pressure, rhythm interval sequence extracted | Automated by the needle reader, reviewed by hand | 1 hour per case |
| P7 | Ground truth: three independent expert raters mark systolic and diastolic. Cases where raters disagree by more than 4 mmHg are rejected, not averaged. | Melissa Tully plus two clinical raters | 20 minutes per rater per case |
| P8 | Acceptance and packaging into the case schema with the capture protocol block | Melissa Tully | 30 minutes per case |

Stage P2 is the one most often skipped and the one most worth protecting. Taking two colleagues all the way through to an accepted case, before any volunteer is scheduled, is what stops a capture day producing fourteen unusable recordings because the audio was in the wrong mode or the slate was missed.

## 4.7 Volume and effort

| Quantity | Figure | Basis |
| --- | --- | --- |
| Subjects for release one | 8 | Priorities one to six plus two spares against rejection |
| Takes per subject | 3 | Allows two rejections and still yields a usable case |
| Raw recordings | 24 | Subjects multiplied by takes |
| Expected accepted cases | 14 to 18 | Assuming a rejection rate of one quarter to one third |
| Capture days | 2 | Six subjects per day at forty five minutes each |
| Post production and annotation hours | 48 | Two hours per raw recording |
| Ground truth hours | 14 | Three raters, twenty minutes, per accepted case |
| Total production effort | approximately 90 hours | Excluding rig build, consent work and scheduling |

Ninety hours is not a large number. The programme has never been effort constrained; it has been permission and recruitment constrained, and those two things run on calendars rather than on hours. That is why P3 starts on day one and runs in parallel with everything else.

## 4.8 Asset specification

Filename       MBPT_<caseId>_<take>_<stream>.<ext>

example      MBPT_C014_T02_gauge.mp4

MBPT_C014_T02_audio.wav

MBPT_C014_T02_pressure.csv

Video          H.264, 1920x1080, 60 fps, constant frame rate,

manual exposure / focus / white balance, no stabilisation

Audio          WAV, PCM 16 bit, 16 kHz or above, mono, no processing,

no normalisation, no noise reduction of any kind

Pressure log   CSV, columns: t_ms, pressure_mmhg, sensor_id

Every case carries a capture_protocol record, and a case without one is

not accepted into the library:

"capture_protocol": {

"device": "<stethoscope make and model>",

"mode": "bell",

"gain": "fixed_setting_4",

"sample_rate_hz": 16000,

"camera": "<make, model, lens>",

"frame_rate": 60,

"exposure_locked": true,

"reference_sensor_id": "REF-01",

"sync_method": "led_tone_slate",

"consent_record": "<reference>",

"captured_on": "2026-09-xx",

"operator": "<initials>"

}

## 4.9 Acceptance criteria for a Class A case

- Three expert raters agree on systolic and on diastolic within 4 millimetres of mercury. Disagreement beyond that rejects the case rather than averaging it, because a case whose correct answer experts cannot agree on cannot fairly score a learner.

- The needle is readable by the automated reader across the entire descent, with no frame where confidence falls below threshold for more than 200 milliseconds.

- No exposure or focus shift across the take, verified by frame differencing on a static region of the dial face.

- No audio clipping and no audible speech.

- Deflation rate within 2 to 3 millimetres of mercury per second for at least eighty percent of the descent.

- Synchronisation offset between audio and video established to within 20 milliseconds.

- A complete capture protocol record.

# 5. Platform and delivery

## 5.1 The requirement

The product must run on the phones learners already carry, which means both Android and iPhone, and it must run on Android XR glasses without being rebuilt. Those two requirements together rule out a phone only native application, and they rule out an application that assumes glasses.

## 5.2 The architecture that satisfies both

| Layer | What it is | Runs where |
| --- | --- | --- |
| Core | Pure logic with no device dependencies: case models, the auscultation engine, the deterministic scorer, mastery rules, trend measures. Written once, tested on a laptop in under a second. | Everywhere. Also runs server side for batch scoring and content validation. |
| Perception interfaces | Declared by the core, implemented per platform: gauge reading, arm measurement, placement tracking. Nothing above this layer knows how a pressure value was obtained. | Interface everywhere, implementation per device. |
| Web application | The product. Progressive web application, installable, works offline, uses the standard browser camera and audio interfaces. | Every phone including iPhone, every tablet, every laptop, and the browser on Android XR devices. |
| Native shell | A thin wrapper around the same web application, for app store distribution and for camera access where a browser cannot reach it. | Google Play, including the store on Android XR devices. Apple App Store if and when wanted. |
| Spatial layer | Added later. Depth, anchors and hand tracking for the glasses specific activities. | Android XR devices only. Absent on phones, which lose nothing because those activities are glasses specific by design. |

Why web first rather than native first A web application reaches every learner on the device already in their pocket, with no installation and no platform exclusion. A native shell around that same application adds store distribution and deeper camera access without a second codebase. Going native first would exclude iPhone learners on day one and force a rewrite to recover them. Going web first costs a small amount of camera capability on glasses, which is recoverable through the shell.

## 5.3 How a reading is actually taken

- The learner points the device camera at the real aneroid dial. The dial is high contrast by design, engineered to be read at a glance, and it is already in frame because the learner is looking at it.

- A needle angle estimator reads the dial thirty to sixty times a second. Two point calibration, taken once per session in about fifteen seconds, maps angle to millimetres of mercury.

- That pressure value indexes the case audio. A cardiac clock advances in real time and fires beat requests; the pressure state decides whether a sound is audible; a beat plays only when both agree.

- The learner marks systolic and diastolic by voice, because both hands are on the equipment. The mark is timestamped at voice onset rather than at recognition, so the timing is carried by onset detection rather than by a language model.

## 5.4 Two numbers that make this tractable

At a correct deflation rate of two to three millimetres of mercury per second, one hundred milliseconds of processing latency corresponds to roughly a quarter of a millimetre of mercury. That is an order of magnitude inside the four millimetre scoring tolerance and below what a human can read off a dial.

So pressure accuracy is not the binding constraint. Perceived audiovisual synchrony is, and that gives a budget of roughly one hundred and twenty milliseconds end to end, which a thirty frame per second pipeline meets comfortably.

One practical consequence: wired earphones, not Bluetooth. Standard Bluetooth adds one hundred to two hundred milliseconds, which barely affects accuracy but sits at the edge of perceptible lag against a needle the learner is watching. The application should detect a Bluetooth output route and warn.

## 5.5 Degraded modes, designed rather than discovered

| Condition | Behaviour |
| --- | --- |
| Dial lost for under 500 milliseconds | Extrapolate at the last observed rate. Continue audio. Say nothing. |
| Dial lost for 500 milliseconds to 3 seconds | Hold the last pressure, suppress new beats, show a small peripheral indicator. No spoken interruption. |
| Dial lost for over 3 seconds | Pause the attempt, mark the interval in the record, prompt the learner to bring the dial into view. |
| Confidence below threshold while visible | Treat as lost. A wrong pressure is worse than none, because it fires beats at wrong values and teaches a false association. |
| Implausible rise with no squeeze detected | Reject the sample. Physical systems do not spontaneously repressurise. |
| No camera, or camera reading unavailable | Fall back to paced shadowing: the authored deflation ramp drives the same engine and the learner marks against it. A legitimate pacing and perception exercise, labelled as such, not a failure state. |

# 6. Assessment and scoring

## 6.1 Four principles

- Deterministic. The same attempt record produces the same score every time, on every device, forever. This is what makes a score defensible to a programme director and auditable to an accreditor.

- Thresholds live in configuration, never in code. Changing a tolerance is a data change with a version number.

- Generative models phrase feedback. They never grade. A model that both grades and explains will eventually explain a grade it invented, so the model receives a score it cannot change and is asked only to word it. Any response containing a number absent from the scorer output is discarded and a templated fallback shown.

- Trends are scored separately from attempts. Rounding bias is meaningless in one attempt and diagnostic across ten.

## 6.2 The thresholds that govern scoring

| Parameter | Value | Note |
| --- | --- | --- |
| Marking tolerance | plus or minus 4 mmHg | Matches the case acceptance threshold deliberately |
| Deflation rate band | 2 to 3 mmHg per second | The most common technique fault by a wide margin |
| Inflation margin | at least 30 mmHg above the palpated estimate | Underinflation is what turns a gap into a misdiagnosis |
| Cuff edge placement | 2 to 3 cm above the antecubital fossa |  |
| Cuff bladder fit | DECISION REQUIRED | The current professional standard and the widely taught legacy rule differ. See section 9. |
| Sequence mastery | score at or above 90 on three consecutive attempts |  |
| Terminal digit window | 10 attempts | Below the window size a trend measure has no meaning |

## 6.3 What an attempt record must carry

Beyond the obvious. Each of these exists because an objective cannot be scored without it, and each has been missed in earlier planning.

| Field | Needed by | Why |
| --- | --- | --- |
| scaffold_state | Every Domain B, C and D mastery rule | Scaffolded success is progress, not competence. Without this field the two are indistinguishable. |
| prompted flag on spoken responses | E4 escalation | A learner who names the escalation only when asked whether they should escalate has not met the objective. |
| surface | Coverage reporting | Which device the attempt was made on determines which objectives were assessable at all. |
| rate_excursions | C3 | A learner averaging 2.5 by alternating between 0.5 and 5 has not learned control. |
| reinflation_events | C5 | Re-inflating on a partly deflated cuff is a technique fault the mean rate hides. |
| confidence and tracking gaps | D2, D3 | An attempt where the dial was lost during the marking window cannot be scored for accuracy, and must be flagged rather than scored badly. |

# 7. The way forward

## 7.1 Three workstreams, running in parallel

The critical insight about sequencing is that these do not queue behind each other. The media programme is calendar bound, the software is effort bound, and the curriculum is already largely done. Running them serially would waste the autumn.

| Workstream | Bound by | Starts | First milestone |
| --- | --- | --- | --- |
| Media production | Calendar: permissions, recruitment, scheduling | Immediately. Permission work on day one. | Pilot capture accepted end to end, week 3 |
| Software | Effort: engineering hours available | Week 1, with the needle reader | Needle reader validated against reference pressure, week 3 |
| Curriculum and content | Dependent on the needle reader for case annotation | Week 3 | Ten existing cases migrated, week 6 |

## 7.2 The first thing to build, and why it cannot be wasted

Build the needle estimator as a standalone offline tool before any application exists. Point it at recorded video of dials and compare its output against a reference pressure sensor.

It is the decisive technical gate, because the closed loop depends on it. But it is also the content annotation tool, because tagging each extracted heartbeat with the pressure at which it occurred requires reading pressure off your own recordings. And it is the capture rig validator, because it is what verifies a take is readable before a volunteer is thanked and sent home.

One tool, three jobs If the needle estimator hits accuracy, the product has a closed loop. If it is marginal, the product falls back to paced shadowing and the estimator still annotates the entire case library. If it fails outright, it has still validated the capture rig. There is no outcome in which building it first is the wrong move.

## 7.3 Milestones

| Milestone | Depends on | Evidence it is met |
| --- | --- | --- |
| M1, needle estimator validated | Recorded dial video plus a reference sensor | Within 2 mmHg on 95 percent of frames, across three gauges and three lighting conditions. Error distribution reported, not just the mean. |
| M2, capture pipeline proven | M1, rig built, protocol frozen | Two colleague recordings taken all the way through to accepted cases |
| M3, permissions in place | Site identified, consent form approved | Written institutional permission and a not human subjects research determination |
| M4, case library migrated | M1 | Ten existing cases in the current schema with beat libraries and rhythm sequences |
| M5, core loop working | M1, M4 | A learner deflates a real cuff and hears the correct case sounds at the correct pressures |
| M6, curriculum implemented | M5 | Fifty two first release activities, crosswalk showing zero coverage faults |
| M7, pilot ready | M6, instructor pack, dry run | Instructors have run a full session on the real equipment in the real room |
| M8, pathology cases accepted | M2, M3, capture sessions | Priorities one to four accepted. Unlocks the five held back objectives. |

Milestones M1 to M7 are the first release. M8 is the second and it is the one that completes the product's central claim, which is why the media workstream starts on day one even though it delivers last.

# 8. Risks

| ID | Risk | Rating | Response |
| --- | --- | --- | --- |
| R1 | Clinical recordings are not acquired, so the product never covers pathology and the human variation claim stays partly unevidenced. | High likelihood, high impact | Media production is a named workstream with its own owner and its own calendar, starting on day one. Community volunteer recruitment rather than hospital patients removes a layer of gatekeeping. Section 4 is the mitigation in full. |
| R2 | Needle reading does not reach the required accuracy in classroom conditions. | Medium, very high impact | Proven offline before any application code exists. Paced shadowing is a complete shippable product, and the estimator still serves annotation. |
| R3 | Capture protocol drift makes the library internally incomparable. | Medium, high impact | One device, one mode, one gain, one rig. Recorded in every case and audited at acceptance. Locked manual exposure is the single most important setting. |
| R4 | Consent obtained does not cover commercial educational distribution, discovered after capture. | Medium, high impact | Consent form drafted and reviewed before the pilot capture, not before the first volunteer session. Naming commercial distribution explicitly is the whole point. |
| R5 | Paediatric recording is treated as one subject among eight and blocks the release. | Medium, medium impact | Paediatric is deliberately last in priority and is not in the first release scope. Treat it as a separate mini programme. |
| R6 | Engineering capacity is insufficient for the activity count in the timeframe. | Medium, high impact | Contract the undifferentiated interface work. Keep the needle estimator, the scorer and the content pipeline in house, because those are the durable assets. |
| R7 | The cuff sizing rubric encodes a standard the partner institution does not teach. | Medium, medium impact | Confirm in writing before the rubric freezes. The configuration already supports either. |
| R8 | Learners use Bluetooth earphones and experience audio lagging the needle. | High, low impact | Detect the output route and warn. Consider refusing to score attempts made on a Bluetooth route. |
| R9 | Scope creep from spatial features into the first release. | Medium, medium impact | The first release is phones. Glasses specific activities are identified in the crosswalk and excluded by release tag. |

# 9. Open questions

| Question | Owner | Needed by | Consequence of not answering |
| --- | --- | --- | --- |
| Which cuff bladder fit standard does the rubric encode: the current professional statement, or the widely taught legacy rule? | Nisha Patel with the partner institution | Before the rubric freezes | Objective B2 is unscoreable, and B2 carries the acceptance criteria. |
| Which site hosts the recording, and is a not human subjects research determination available there? | Melissa Tully | Week 1 | The entire media programme, and with it five objectives. |
| What is the scope of practice boundary for escalation, per cohort? A patient care technician and a practical nursing student owe different actions on a crisis value. | Nisha Patel | Before Domain E content is authored | One rubric cannot serve both cohorts without a scope parameter in the case data. |
| Do partner programmes supply devices at the station, or is bring your own device assumed? | Nisha Patel | Before the pilot is scoped | Determines whether the reach architecture in section 5 is sufficient. |
| Which gauge models are actually in use at the partner institutions? | Nisha Patel | Before milestone M1 completes | The needle reader must be validated against the dials learners will really point a camera at. |
| Who are the three expert raters for ground truth, and are they available for the volume in section 4.7? | Melissa Tully | Before capture sessions | Cases cannot be accepted into the library without them. |
| Is the pilot date set? | Melissa Tully | Immediately | Every milestone in section 7.3 is currently relative rather than dated. |

## 9.1 Records to open

- A capture protocol specification, frozen, printed, and kept at the rig. It is referenced by every case record and it must not be an email thread.

- A consent form and media release naming commercial educational distribution, reviewed by counsel.

- An architecture decision record for the web first plus native shell approach, recording the reach argument.

- An architecture decision record for optical gauge reading, recording the latency reasoning and the validation result.

- A take log template, which is the only place subject identifiers ever appear and which is never shipped.

End of document.